## Slide 1: Transmission and Reflection Spectroscopy: Calculations
James Haggerty
Department of Physics
Oregon State University
February 16, 2017

## Slide 2: Motivation
To explain how to process transmission and reflection data use the optical calculation spreadsheets.
To understand how the quality checking and calculation spreadsheets work

## Slide 3: Processing Flow Chart: Optical Plots
Obtain T & R Data
Which Spectrometer (OO or GS)?
Ocean Optics: Single Detector
Ocean Optics: Both Detectors
Grating Spectrometer: Both Detectors
Which Detector (Si or InGaAs)?
Grating Spectrometer: Single Detector
Use Single Range Processing Flow
Use Multi-Range Processing Flow

## Slide 4: Processing Flow Chart: Optical Calculations
Optical Plots
Index_Calc Spreadsheet
Tauc_Gap Calculator
Diff_Gap Calculator

## Slide 5: Processing Flow Chart: Index Calc
Copy Raw data from Optical Plots Spreadsheet
This is a manual copy: Make sure to use the paste special option for pasting only values. If concatenating data make sure to paste them in the correct order with increasing wavelength with no overlap between the two datasets.
Quality check T and R plots to make sure the data was copied correct
Defining fringes and best guess at thickness
This is to initialize the spreadsheet. All these values will be adjusted to obtain the best fit for the data
Processing data to obtain best fit
Finding the interference fringe value of the first fringe
Adjusting the thickness to fit the simple index to the fringe data
Adjust the position of the fringes?
Fit the Sellmeier index to the simple index
Adjust the thickness to fit Sellmeier index to fringe data.

## Slide 6: Copy Data from Optical Plots to Index_Calc
Copy the data manually from either the Optical Plots spreadsheet or the Make_Int_Step spreadsheet
  Make sure to use the paste special option of only pasting values
  Make sure to apply the data filters to standardize all the plots.
Also don’t forget to go to the Info tab and fill out the demographic information
  Notice the drop down menu option that allows you to process either α or αd.

## Slide 7: Index Model for substrate
Before we get into any of the data processing we first need to make sure we are using the correct index model for our substrate
  Luckily, there is a database which contains index data for most of the substrates we use
    http://refractiveindex.info/
  To account for the different types of substrate information we obtain from this database we have 3 options for determining the index of refraction for our substrate
    0=Sellmeier equation
    1=Cauchy equation
    2=Experimentally measured T and R data
  To use options 0 or 1 you must have the coefficients for the substrate
  If the database doesn’t contain these parameters we can back calculate the index from our measured data
    Entered into the substrate – data tab

## Slide 8: Using reference data to calculate index of refraction for substrates
Since some substrates do not have fit parameters for them (e.g. Eagle XG glass) and the index values given do not cover the entire range of our film data we can use our experimentally measured T and R data for our substrates to create a model for the index within the range for our data.
  This is accomplished using the multiple interference equations from Optics by Hecht.
  For EXG glass this produces a 0.03 difference between the reference data and the calculated data

## Slide 9: Define Interference Fringes and Film Thickness
Go to the calc-index tab
  On the plot of R find the wavelength associated with the minima and maxima of the interference fringes
    Enter them into the column directly above the plot of R
    Starting from the largest wavelength and going to the shortest wavelength.
  Next fill in your initial estimate for the thickness of the film (on LHS)
    This does not need to be perfect as we will adjust the thickness to fit the data.
  Lastly, make sure to fill out the fit range min and max values.
Notes: If you place the cursor over the position of the fringe on the plot of R it will give you the wavelength value for that fringe.

## Slide 10: Defining the appropriate m₀ value
Next we need to define the interference order value (m₀) for the first fringe identified
  integer
    When R is max
  half integer
    When R is a min
Getting m₀ can be a difficult task but it is a very important parameter

## Slide 11: Choosing the correct m₀ value
m₀ is too low on gph-index
m₀ is too high on gph-index
The n:simple should be a monotonically decreasing function with increasing wavelength.
The n:simple function should never go below the index of the substrate
  Unless you have good reason
  In which case this model is not correct for your material stack.

## Slide 12: Choosing the correct m₀ value
Effect of bad thickness guess
m₀ just right on gph-index
Even if your m₀ value is correct if your first guess for the thickness is too far off you will get bad values for the index
  In the cases to the right
    If thickness is too high the index values become unreasonably low
    If thickness is too low the index values become unreasonably high
  You must keep in mind the index values your material should have
    Remember that it doesn’t have to be perfect but the closer the guess the more likely you are to get the correct results.

## Slide 13: Adjust thickness to fit simple n to fringe data
Now we can begin to process the data, starting with Macro button 1)
  This button adjusts the thickness (d) until the simple index function fits the fringe data
  The top plot is before the thickness adjustment and the bottom one is after the thickness adjustment

## Slide 14: Fringe Adjustment: Example 1, Oxides
Next we can look at the before (top) and the after (bottom) thickness adjustment to see how our simple model fits the data
  The first thing we notice is that the fringes we defined don’t match up with the fringes in the data as well as we might like
  This is where macro button 1.5) comes to play
    This is not always needed but can be quite helpful for some of the more difficult datasets.

## Slide 15: Fringe Adjustment : Example 1, Oxides
Index of Refraction Before
Reflectance Before
Index of Refraction After
Reflectance After
Notes: In this case it did not make a huge difference (2 nm difference in thickness) but now the fringes line up with the data better and the trend in the index has not changed significantly.

## Slide 16: Fringe Adjustment : Example 1, Oxides
Index of Refraction without
Reflectance without
Index of Refraction with
Reflectance with
Notes: In this case it did not make a huge difference (2 nm difference in thickness) but now the fringes line up with the data better and the trend in the index has not changed significantly.

## Slide 17: Fringe Adjustment: Example 2, Selenides
Index of Refraction Before
Reflectance Before
Index of Refraction After
Reflectance After
Notes: In this case it did not make a huge difference (2 nm difference in thickness) but now the fringes line up with the data better and the trend in the index has not changed significantly.

## Slide 18: Fringe Adjustment : Example 2, Selenides
Index of Refraction without
Reflectance without
Index of Refraction with
Reflectance with

## Slide 19: Fitting Sellmeier index to fringe data
Once we obtain the best fit possible with the simple index we then need to fit the Sellmeier equation to the fringe data (macro buttons 2) and 3))
  The simple fit gives us a better starting point for fitting the Sellmeier index
  Something to keep in mind for the Sellmeier fit is the asymptote named
    This value tends to creep up into the range for the data and produce an unphysical result (if then R can become unrealistic)
    can however be constrained with the Constrain parameter which restricts the value of to the minimum fit range value minus the constrain value.

## Slide 20: Assessing fit criteria
There are two fit criteria for determining if the results are good
  First, your judgement is an important criteria as the program doesn’t know what is physical or not it’s only trying to find a minimum
    Use your knowledge of the data and what should be reasonable
  Second, the is the minimization parameter for the program
    It’s cell is color coded to show what should be a decent fit
      Green is the best fit:
      Yellow is acceptable: 1< <2
      Red is a bad fit: >2

## Slide 21: Assessing Absorption
The very last thing we need to check in the index_calc spreadsheet is the absorption plot.
  Does this plot make sense based on what you would expect from your material?
Now that we have the absorption for our material we can calculate the band gap.

## Slide 22: Processing Flow Chart: Tauc Gap
Copy T, R, and α data from Index Calc Spreadsheet
This is an automatic copy similar to the one used in the Optical Plot Spreadsheet
Quality check T, R, and α plots to make sure the data was copied correct
Defining Linearized regions for gap calculation
Direct gaps and indirect gaps, allowed gaps vs forbidden gaps
Main linearized regions and baseline corrections for sub gap absorption

## Slide 23: Copy T, R, and α data
This copy function works exactly the same as in the optical plots spreadsheets, except for a few different parameters
  First, T and R, and α don’t have to be on the same Tab in the linked file
    So we need to define the Tab’s that they are on, one for T and R and one for α.
    The default parameters here are set to link into the index_calc spreadsheet but can be changed to link into the optical plots spreadsheets as well.
  Second there are a couple different drop down menus for plot formatting
    “Order to Plot α” allows you to scale the axes that α is plotted on making the plot more visually friendly
    “α units” gives you the option of choosing between absolute units or arbitrary units
      Absolute units is the default option
      Choosing the blank option in the drop down gives you the additional option of which type of relative absorption you wish to analyze
        For a quick analysis of thin film absorption from the T/(1-R) data you would analyze αd (absorption times film thickness)
        For an analyses of diffuse absorption you would analyze α/S (absorption divided by scattering factor)

## Slide 24: Absorption Plots
Included in this spreadsheet are formatted (“pretty”) plots of α that include T and R data.
The additional Demographic parameter of “Data Set, Short Name” on the previous slide can be used to give these plots names specific to the film that aren’t included in the title.
  Be careful if you change this parameter though, and make sure to follow the notes in the spreadsheet when you change this parameter.

## Slide 25: Defining Linearized Region: Gap Calculation
The first thing we need to do here is to define which set of data we want to process on the Gap Range Select Tab.
  Allowed or Forbidden Gaps
    The drop down menu here will allow you to change all of the plots from Allowed Gaps to Forbidden gaps with a single option.
    All plots and axis information are updated to reflect the dataset you are analyzing

## Slide 26: Defining Linearized Region: Gap Calculation
Using the Tauc plots given on the Gap Range Select Tab define the min and max values for the Direct and Indirect linear regions.
  There is a bit of choice here and so we need a way to assess the “linearness” of the regions we choose
  Also we need to account for any baseline sub gap absorption
    There isn’t much in this dataset but we can see that the indirect plot doesn’t quite reach 0 so we will account for that with the baseline range select

## Slide 27: Defining Linearized Region: Gap Calculation
Once you define the linearized range for the gaps and the baseline correction ranges press the Range Select and Calculate Gap button
When this macro finishes a Tauc plot is displayed with the linear functions for the regions you defined
  Notice that our first guess for the linear baseline doesn’t quite capture the character of the baseline
    If this happens we need to go back the Gap Range Select Tab and choose a different region for the baseline until we get a result that looks appropriate

## Slide 28: Defining Linearized Region: Gap Calculation
Once the baseline gap correction is satisfactorily representing the baseline absorption we can now assess our gap calculations
  Back on the Gap Range Select Tab there were R² values that were originally giving an error
  Now that the gaps are calculated these values are now available
    This is the R² value used to assess how linear our regions actually are
      The closer to 1 these values are the more linear our region.
        One 9 after the decimal and you might want to choose a different region
        Two 9’s after the decimal is a good fit
        Three 9’s is better
        A value of exactly 1 is extremely unlikely

## Slide 29: Defining Linearized Region: Gap Calculation
Once we determined that our linear fit is acceptable we are done.
The Tauc plot now shows all the information about the band gap for our material

## Slide 30: Processing Flow Chart: Diff Gap
Copy T, R, and α data from Index Calc Spreadsheet
This is an automatic copy similar to the one used in the Optical Plot Spreadsheet
Quality check T, R, and α plots to make sure the data was copied correct
Alternate Band Gap Calculations
Discrete derivatives and smoothing
Gaussian fitting for gap determination
Determining Gap Type
Region of linearization determination

## Slide 31: Copy and smoothing parameters
The demographic information here is exactly the same as that found in the Tauc Gap calculator except for one parameter # λ points to smooth (odd) and Gap Estimate
  The # λ points is the smoothing parameter needed to smooth discrete data when taking a derivative.
  This parameter has to be odd because it is using a centered average function with an even number of cells above and below the center point,
    So for a value of 7 there will be 3 wavelengths above and 3 wavelengths below which are averaged together
      For example at 250nm, wavelengths 247, 248, 249, 250, 251, 252, and 253 are averaged together to smooth the data, and assigned the value of 250nm
    We will check the validity of this smoothing value later but for now the default 7 is a good place to start.

## Slide 32: Motivation for Diff Calculation
The reason this spreadsheet exists is due to the fact that when using the Tauc method of band gap calculations we must assume a gap type. In this method we don’t have to assume a gap type at all.
  Both methods rely on the theoretical model for absorption from DOS calculations.

## Slide 33: Motivation for Diff Calculation
  is absorption
  is a constant
  is energy and is the band gap in energy units
  is the constant associated with the gap type
    for a direct allowed band gap
    for a indirect allowed band gap
    for a direct forbidden band gap
    for a indirect forbidden band gap

## Slide 34: Motivation for Diff Calculation
  is the constant associated with the gap type
    Notice that in the Tauc analysis we assumed the band gap was either indirect or direct and allowed and plotted or respectively. Consider this example
    This produces a linear function near the band gap if our material is actually a direct gap material, but we had to assume that from the beginning!

## Slide 35: Motivation for Diff Calculation
Let us take a different approach, consider the function
    First thing to notice about this calculation is that if we plot vs we should have linear function with slope , which tells us the gap type
    Second thing is, what if we consider the derivative of this expression with respect to energy
  This time we get an asymptotic function centered at the band gap.
  Thus we should be able to obtain a value for the band gap without assuming the type.

## Slide 36: Motivation for Diff Calculation
One thing that must be mentioned is that our data are discrete and not continuous as the derivative we are taking suggests and so we must take that into account in our calculations.
  First, discrete data are inherently noisy and thus we apply a smoothing function (as described previously) to make the differentiation cleaner
  Second, the derivative we take is a symmetric discrete derivative which is more accurate.
      Here b and a correspond to before and after, as in the point just before the point of interest and the point just after the point of interest.

## Slide 37: Derivative Gap Calculation
Once the data is copied over to the diff_calc spreadsheet we should be taken to the Band Gap Calc Tap
  First we should see a plot similar to the one shown here
  Second we will do a Gaussian fit to the derivative curve, given by

## Slide 38: Derivative Gap Calculation: Smoothing
Once we have adjusted the adjusted the initial values but before we fit our data we need to look at the plot on the Alpha eV Smoothing Tab
  First, we need to make sure the smoothed data matches the raw
  Second, we need to make sure that the difference between them is relatively small.
  If either of these looks unreasonable we should return to the Info Tab and adjust the smoothing parameter either up or down in odd numbered increments.
    Be careful when you adjust this parameter as it has a tendency of making the derivative on the previous slide noisier.

## Slide 39: Derivative Gap Calculation: Smoothing
Max Smoothing (15)
Default Smoothing (7)
Min Smoothing (3)
The Main thing to notice with these images is that if we apply too much smoothing we eventually start to take some of the gap information out of our spectrum
  We’ll see the effect of this in the next slide

## Slide 40: Derivative Gap Calculation: Smoothing
Max Smoothing (15)
Default Smoothing (7)
Min Smoothing (3)
While the highest smoothing value has produced the most smooth derivative spectrum, it has also shifted our gap value.
Also, the minimum smoothing appears to produce the same gap amount but with much noisier data.
  More smoothing may be needed on films with even noisier data.

## Slide 41: Derivative Gap Calculation: Gap Type
Now that the gap for our film is calculated we can move on to determining the gap type on the Gap Range Select Tab
  From the plot of vs we need to define the linear region
    This works exactly the same as on the Tauc gap calculator but now the value are a bit weird and so when we read them off the plot they need to be converted back to nm values, for the range select.

## Slide 42: Derivative Gap Calculation: Gap Type
Now here we run into a problem…
  Where do we define the linear range for the gap type?
  If we use the data closest to the gap we just calculated we get and not a terrible linear fit

## Slide 43: Derivative Gap Calculation: Gap Type
Now here we run into a problem…
  Where do we define the linear range for the gap type?
  If we use the data just above the gap we just calculated we get 46 and a bit better fit
  Closer to what we would expect for a direct gap material, but based on the crystal structure of this film, it should not have a direct gap

## Slide 44: Derivative Gap Calculation: Gap Type
Now here we run into a problem…
  Where do we define the linear range for the gap type?
  If we use the data range with the steepest slope we get and an almost perfect linear fit
  Notice that this range is almost exactly the same as the range we used in the Tauc analysis for a direct gap.

## Slide 45: Derivative Gap Calculation: Gap Type
Now one thing we need to discuss is the fact that we got for the most linear section in the last slide
  This value doesn’t correspond to any of the values we would expect for a crystalline material
    This is okay, because we know (from other analysis) that this material is not a single crystal and thus it isn’t perfectly crystalline
    I suspect that most materials would not actually have theoretically expected values for , however this value is still reasonable because it lies somewhere between the values for an indirect () and direct transition (

## Slide 46: Optical Band Gap Calculations
Tauc Gap Calculator
Derivative Gap Calculator