## Slide 1: Transmission and Reflection Spectroscopy: Optical Plots
James Haggerty
Department of Physics
Oregon State University
February 16, 2017

## Slide 2: Motivation
To explain how to process transmission and reflection data use the optical calculation spreadsheets.
To understand how the quality checking and calculation spreadsheets work

## Slide 3: Processing Flow Chart: Start
Obtain T & R Data
Which Spectrometer (OO or GS)?
Ocean Optics: Single Detector
Ocean Optics: Both Detectors
Grating Spectrometer: Both Detectors
Which Detector (Si or InGaAs)?
Grating Spectrometer: Single Detector
Use Single Range Processing Flow
Use Multi-Range Processing Flow

## Slide 4: Processing Flow Chart: Single Range
Copy Raw data from spectrometer text file to Single Range Optical Plots Spreadsheet
Both files (spectrometer file and plot spreadsheet) need to be open for the copy
Quality check data for spurious T and R values and investigate Raw intensity spectra for likely causes.
Ocean Optics or Grating Spectrometer?
For Ocean Optics: Take final T, R, and λ data and copy over to Make_Int_Step spreadsheet to either reduce or interpolate data to interger steps in wavelength
For Grating Spectrometer: Proceed to next step
Copy T, R, and λ data into the Calc_index spreadsheet
Once at this step all data processing becomes the same.

## Slide 5: Raw Data File
Ocean Optics Data File
Grating Spectrometer Data File
The data output style for each spectrometer is slightly different, but how they calculate T and R are essentially the same, where
  TS is the power measured for the light through the sample
  PBG is the power measured by the detector under dark conditions
  PR is the power measured without the sample present (typically called the lamp)
  RS is the power measured for the light reflecting off the sample.
  Rref is the reflectivity of the reference used
    For instance in the OO spectrometer we use a mirror as reference and we need to correct for the fact that the mirror is not 100% reflective for all wavelengths.
    However, for the grating spectrometer no reference is used so Rref=1

## Slide 6: Copy Optical Data for Quality Checking
Spectrometer Data File
Single Range Optical Plots Spreadsheet
The first thing to note about the optical plots spreadsheet (and this is true for all subsequent spreadsheets) is that if a cell is colored light blue then it is expecting an input from the user.
  In this case the input information is demographic
Notes: Both of these files must be open in order to copy the data from one to the other
Current version is named: Optical_Plots_SingleRange-V05.xlsm

## Slide 7: Single Range Optical Plots Spreadsheet
The tab that should come up first when the optical plots spreadsheet is opened is the “Info” tab.
Make sure to fill out all demographic information
  Global Information and Data
  These values are used for plot labeling and data selection
Under Data File Information
  Make sure that the file name of the file which contains the data you want to copy is open and spelled correctly in the file name box (Cell I2)
  The cell references use the following format “Beginning Row”-”Beginning Column”
    The Typical Values sections contains the values for the cell references from each spectrometer file (OO or GS) corresponding to data collected on a single sample for a single measurement
    They are color coded for easy reference.

## Slide 8: Single Range Optical Plots Spreadsheet
Once all demographic information is filled in, notice the Gray buttons on the bottom LHS of the screen
  These buttons are used to do exactly what they are labeled to do
First you want to press the “Copy Data from Linked File” button
  This will copy the data from the spectrometer file and paste it into the optical plots spreadsheet

## Slide 9: Single Range Optical Plots Spreadsheet
Once the data is copied from the spectrometer file the Macro should bring you to the “Raw R” tab which contains a plot like the one shown on the right
  Things to look for in this plot
    Spikes in the lamp spectrum
      Note the wavelength as this could be a possible source of noise
      The spike at 656 nm in the lamp tends to be problematic for this spectrometer
    The order of the spectra
      Lamp should always be higher than Reflection which should be higher than Background
Repeat this QC procedure on the Raw T tab as well
  Make note of any locations where Reflection and Background are similar in intensity as this will be the high absorption region and could contain significant noise.

## Slide 10: Single Range Optical Plots Spreadsheet
Once the Raw data plots are checked for possible noise or data collection issues we can then use the second button “Select the desired range of data”
  Make sure that Plot Range Min and Plot Range Max cells are filled with values appropriate for your data
    These values can be used to select off noisy ends of your spectra

## Slide 11: Clipping off Noisy Data
Once the “Select the desired range of data” button is pushed a plot of TR nm is displayed
  I’ve labeled the noisy ends of the data which we can now go back and change our range and remove the noisy sections
  We want to do this because we don’t want to process noise, just signal.
Noisy Ends

## Slide 12: Clipping off Noisy Data
Here we changed the Min and Max values and then re-pressed the “Select the desired range of data” button
  Notice that the noisy ends no longer appear in our data
  This procedure can be repeated as many times as needed until the data contains minimal noise on the ends.

## Slide 13: Quality Checking Optical Data: T and R
Now we need to check the optical data for additional noise
  First notice the small spike in T around 656 nm
    This is a lamp artifact that the spectrometer could not resolve

## Slide 14: Quality Checking Optical Data: T/(1-R) and 1-T-R
Now we need to check the optical data for additional noise
  First notice that small spike in T around 656 nm propagates through our data
  Also notice around 440 nm the wiggles in the data
    This corresponds to an alignment issue in the OO spectrometer and is very difficult to remove entirely.

## Slide 15: Quality Checking Optical Data: T/(1-R) and 1-T-R
Now we need to check the optical data for additional quality
  First thing we want to note about the quality is that none of the fringes from T and R appear in T/(1-R)
    This corresponds to a sample alignment problem
  The second thing we want to note is that neither T/(1-R) or 1-T-R goes above 1
    This is typically due to a bad measurement where the sample measurement goes above the lamp

## Slide 16: Make OO data Integer Steps
Now if the data was collected on the OO spectrometer then we need to reduce the data to integer steps to make it easier to process.
  The two spectrometers in the OO system contain different numbers of data points at different wavelength steps
    This can cause a problem if you are concatenating data together for a full spectrum measurement
    For consistency and ease of processing it is best to just make the OO data integer steps
On the “Range Select” Tab is the data corresponding to the plots we just QC’d
The Make_Int_Step spreadsheet contains the tabs shown in the image below
  If your data is UV-Vis then you will paste your data into the UV-Vis data tab
  If your data is IR then you will past it into the IR Data tab
  Make sure when you paste the data over that you use the paste special option and only paste numbers
    The formula’s in the optical plots spreadsheet will not make any sense in this spreadsheet.
Notes: OO_Make_Int_Step_UV-IR-V01 uses an averaging algorithm to average all the extra points from the HR400 detector that correspond to the same wavelength. For instance if 250.1 and 250.2 were in the same data set they would be averaged and assigned the new value of 250.
OO_Make_Int_Step_UV-IR-V01 uses an averaging algorithm to find the mid point between two point in the NIR256 detector and repeats that until all integer wavelengths are accounted for. For instance if 800and 805 were the two sequential data points output by the detector then the average 802.5 would be added. Then 800 and 802.5 would be averaged to give 801 and so on until all the wavelengths are accounted for.

## Slide 17: Make OO data Integer Steps
Once the data are copied over switch to the Final UV-Vis Reduced tab
  Notice the yellow box at the top
    Cells which are this color correspond to drop down menus
    In this case the drop down menu is a filter which needs to be applied to remove the zero values.

## Slide 18: Make OO data Integer Steps
The last thing we need to do before moving on to the calculations is to make sure that the interpolated/reduced data does not diverge from the original data
  Here we can see that if anything we reduced the noise but overall the character of the data was maintained.

## Slide 19: Processing Flow Chart: Multi-Range
Copy Raw data from spectrometer text file to Multi-Range Optical Plots Spreadsheet
All 3 files (both spectrometer files and the optical plots spreadsheet) need to be open for the copy
Quality check data for spurious T and R values and investigate Raw intensity spectra for likely causes.
Ocean Optics or Grating Spectrometer?
For Ocean Optics: Take final T, R, and λ data and copy over to Make_Int_Step spreadsheet to either reduce or interpolate data to interger steps in wavelength. Do this for both detectors as they have different resolutions and need to be made to have the same step size for processing in the next step
For Grating Spectrometer: Proceed to next step
Copy T, R, and λ data into the Calc_index spreadsheet
Once at this step all data processing becomes the same.
Make sure the spectra overlap each other
Use scaling values on the data collected for one of the detectors (i.e. Si) for both T and R to find the best transition point between detectors.
Keep in mind that if this scaling factor gets to large (>10%) you will have to retake the data and fix any alignment issues.

## Slide 20: Multi-Range Optical Plots Spreadsheet
The copy function for this spreadsheet is exactly the same as the single range spreadsheet, except now we have two input datasets to define and work with.
  Make sure to fill out all demographic information for the type of data being analyzed
  Make sure that both spectrometer files are named correctly and are open with the optical plots spreadsheet.
  Once every thing is filled out and the values for the location of the data are set for the particular spectrometer you are using are set, use the Copy Data Set button for each Set 1 and Set 2.
  Quality check the data exactly the same as you would for the single range data, but now you have to do it twice.

## Slide 21: Multi-Range Optical Plots Spreadsheet
Once the data are copied to the optical plots spreadsheet and the data ranges are defined for each dataset, go ahead and push the “Select Defined Range of Data” button
  When this macro finishes you should se a plot of both of your datasets on the same axis
  Notice that in this measurement there are some interesting things going on that we should check out before moving forward with the data processing.

## Slide 22: Multi-Range Optical Plots Spreadsheet
The first interesting feature is that the dataset overlap but they don’t quite match.
  This problem can be due to multiple factors but most likely the cause is due to the different detectors used and their different resolutions
    Our Si detector has a rather larger surface area, while our InGaAs detector has a very small surface area
      Thus their abilities to collect light are very different

## Slide 23: Multi-Range Optical Plots Spreadsheet
The first interesting feature is that the dataset overlap but they don’t quite match.
  To fix this problem, notice the additional parameters on the Info Tab (that didn’t exist on the single range plot spreadsheets)
    T Scaling Factor
    R Scaling Factor
    These are multiplicative/correction factors for the different detectors
      Apply the correction factor to only one detectors data
      The T and R correction factors do not have to be the same as the data was collected under different condition
      Try to find a correction factor that gets the datasets to match that is less than 10% (0.1)
        If larger than 10% it is a good idea to recollect the data and fix any inconsistencies or alignment issues.

## Slide 24: Multi-Range Optical Plots Spreadsheet
The first interesting feature is that the dataset overlap but they don’t quite match.
  To fix this problem, notice the additional parameters on the Info Tab (that didn’t exist on the single range plot spreadsheets)
    Each time the scaling factor is changed you must apply the Select Defined Range of Data macro button
    Once you decided on your range go back to the info tab and change your Range Min and Max values so the dataset no longer overlap but form a continuous set of data
      Note once you filter the data this way you may wish to revisit the scaling factors so there are no jumps in the full dataset.

## Slide 25: Multi-Range Optical Plots Spreadsheet
The next feature of interest to note in these datasets is that the end points don’t appear to be noisy but instead have interesting features
  We will investigate this by looking at the Raw spectra to see if these ends are a problem that we should remove.

## Slide 26: Multi-Range Optical Plots Spreadsheet
Now looking at the Raw spectra for T and R in the high wavelength range we can see that while the intensities are low there is still enough difference between lamp, sample, and dark to say that this feature is real.

## Slide 27: Multi-Range Optical Plots Spreadsheet
Now looking at the Raw spectra for T and R in the low wavelength range we can see that while the intensities are low there is still enough difference between lamp, sample, and dark to say that this feature is real, up to about 220 nm which is where we will chop the data off.

## Slide 28: Multi-Range Optical Plots Spreadsheet
The last thing to note in this plot is that the IR data is quite noisy.
  We can either live with this noise or do a spot fix to replace the values with the surrounding ones
    If we ignore the noise, most likely we will receive a bunch of errors when processing the data that we can ignore (be careful doing this as some of them may be valid errors)
    If we spot fix these points it is done manually and can be done however you see fit
      The way I typically do this is to find the bad point and do and average between the surrounding points so as to create a smooth transition

## Slide 29: Manually De-spiking the data