{"text_format":"2","workbook":"xls","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":1,"images":0,"macros":0,"external_links":1,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Sheet1","index":0,"n_rows":42,"n_cols":8,"rows_kept":42,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}
["Temp (K)","S (mV/oC)",{"$nulls":4},"RK&TM","2005-03-09"]
[3.75,0.735]
[10,2.044,null,null,null,"Type K thermocouple temperature calibration curve"]
[20,4.185]
[30,6.427,"Reference:","The calibration of thermocouples and thermocouple materials [microform] / G.W. Burns and M.G. Scroger. "]
[40,8.639,null,"QC100 .U565  no.250-35",null,null,"Valley Library"]
[50,10.762]
[60,12.781]
[70,14.701]
[80,16.531]
[90,18.285]
[100,19.972]
[110,21.597]
[120,23.664]
[130,24.673]
[140,26.624]
[150,27.517]
[160,28.851]
[170,30.124]
[180,31.335]
[190,32.482]
[200,33.564]
[210,34.581]
[220,35.531]
[230,36.419]
[240,37.245]
[250,38.01]
[260,38.705]
[270,39.3]
[280,39.78]
[290,40.52]
[300,40.589]
[310,40.918]
[320,41.179]
[330,41.376]
[340,41.499]
[350,41.546]
[360,41.815]
[370,41.412]
[380,41.248]
[390,41.039]
[400,40.804]

{"sheet":"Sheet2","index":1,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[],"flags":["formulas_cached_only"]}